A millifluidic device to apply the bulge test to biofilms

Literature on the bulge test has typically focused on its application to metal sheets, polymer films, and biological skin tissue1,2,11,12,3–10. The delicate nature of biofilms, however, requires a device that applies much smaller pressure differences than these previous use cases.

To adapt the bulge test to biofilms, we constructed a custom millifluidic device suitable for imposing Pa-kPa pressure differences across a freely suspended biofilm, equipped with an optical coherence tomography system to quantify the resulting changes in the shape of the biofilm.

Clear cast acrylic sheets (1/16” thickness, McMaster-Carr) were used as the base of the millifluidic device. Two sets of channels were etched or cut into the acrylic using a laser cutter (Industrial Laser ILS 9.75), and the individual pieces were assembled with epoxy adhesive to form the device in two parts that could be assembled reversibly (Figure 4.1A). One vertical cylindrical hole and horizontal channel allowed “Reservoir 1” to be connected the top face of the device and control the pressure on the top face of the sample (p1 in Figure 2.1B); another vertical cylindrical hole passed through the top part and connected to the horizontal channel in the lower part, allowing “Reservoir 2” to control the pressure below the sample (p2 in Figure 4.1B). The channel above the sample was sealed with a glass cover slip to enable imaging.  For the connections to the top face, thicker acrylic slabs with through holes threaded to accept quick-tube coupling fittings (51525K442, McMaster-Carr) were bonded to the top surface. The reservoirs themselves were syringes of known internal diameters.

Figure 4.1 (A) Schematic of bulge test device, fabricated as two parts that are separated to load a sample in the central chamber at the center of the device; the top and bottom parts seal with vacuum grease. When loaded and sealed, port 1 connects to a reservoir of fluid that is used to control pressure on the top face of the sample (not shown); similarly, port 2 permits control of the pressure on the bottom face of the sample. Gray layers are acrylic; blue layer indicates etched channel; green thin layer is cover slip glass; layers are bonded using epoxy. Horizontal channels are longer than shown. (B) Schematic diagram of central chamber.  Bacterial film sample is supported by two washer-shaped disks. A thin O-ring seals this “sandwich” to the top half of the device.

Sample loading protocol

Prior to mechanical testing, the bacterial films had to be peeled from their polycarbonate substrates. To do this, a polycarbonate filter with bacterial film grown on top was placed into a petri dish containing sterile phosphate buffered saline (PBS, Corning, 1X). A 3 mm diameter biopsy punch (Integra Biosciences) was used to gently make a circular cut through the bacterial film but not the polycarbonate. The sample support disk (copper TEM disk, 1.5 mm diameter circular aperture, 3.05 mm outer diameter, thickness 25 µm, Ted Pella) was then slid in between the bacterial film and polycarbonate from the outer edge of the film and used to gently separate the film from the polycarbonate. Once the edge of the film was freed, the cut 3 mm bacterial film disk freely floated away from its polycarbonate support, and it could be lifted out of the PBS using the support disk and placed into the bottom chamber of the device (pre-filled with PBS). A second support disk was placed on top to sandwich the bacterial film, followed by an o-ring (Precision Associates, Inc).

Optional figure to show the biofilm piece prior to testing.

Prior to sealing the two halves of the device together, a 5 ml syringe was filled with PBS and capped with a 30G needle connected to thin silicone tubing (0.31 mm ID, 0.64 mm OD, HelixMark). This strategy was devised to avoid air bubbles in the viewing window as the device was filled with PBS. The tubing was threaded through into the Luer socket on the top half of the device and placed along the top channel using forceps, such that the tube terminated in the viewing window of the chamber that would eventually be directly above the sample. Following this, the two chamber halves were sealed using a thin layer of vacuum grease. The two chambers were then filled simultaneously with PBS: the upper channel using the tubing connected to the PBS-filled syringe, and the lower channel by way of the associated Luer slip connector, drop by drop, until both connectors were filled, following which the tubing was gently extracted and two reservoirs were attached to the Luer slip connectors on either side. The reservoirs could be syringes of any known internal diameter (ID) with a Luer lock or Luer slip attachment.

Syringes used as reservoirs in the collected datasets were 60 ml syringes (BD Scientific) with ID = 26.72 mm. Where specified, certain samples were tested in HEPES buffer (20 mM HEPES, 115 mM NaCl, 1.2 mM MgCl2, pH 7.0). The apparatus was rinsed with soap and DI water and allowed to dry between uses; separate reservoirs and syringes were used for different buffers.

Principle of operation

As indicated in OPTIONAL FIGURE BELOW, each face of the mounted bacterial film disk is in contact with a different reservoir whose applied hydrostatic pressure is controlled independently by varying the height of PBS in the reservoir. As a precaution against possible changes of the biofilm with time in PBS, experiments were completed within 30 minutes after mounting a specimen. For these experiments, the 60 ml syringes serve as reservoirs with cross-sectional area of 5.61 cm2, such that adding 858 µl of PBS increases the reservoir level by 1.53 mm, which increases the hydrostatic pressure on the corresponding face of the sample by 15.0 Pa. (Refer to figure here)

Figure ???. Schematic of bulge test process. As fluid height increases in the reservoir connected to the chamber underneath the film, the film “inflates” through central aperture (gray box).

A syringe pump was used to impose various pressure-time profiles (Figure 4.2). “Ramp” and “sawtooth” profiles involved PBS addition at a flow rate of 70 ml/min, with a final target volume of 55 ml (ramp) and 2 ml (sawtooth) unless otherwise specified. The flow rate was chosen to be the maximum accommodated by the syringe pump in order to minimize perfusion through the biofilm for the duration of the experiment.

Figure 4.2 Schematic of pressure profiles imposed during bulge test. Stepwise used for 3D measurements (Chapter 4); ramp profile used at varying rates for 2D measurements (Chapter 3), sawtooth profile used to estimate viscoelasticity and fatigue.

Imaging

All OCT imaging was performed with a Thorlabs OCT (GAN210 base unit: 930 nm central wavelength, 6/4.5 µm axial resolution in air/water, 2.9/2.2 mm imaging depth (air/water), OCTP-900 scan head, OCT-LK3-BB scan lens: 36 mm FL, 8 µm lateral resolution). A-Scan/Line Rate was 36 kHz for all measurements (acquisition time = 19 ms). Biofilm thicknesses were calculated estimating a refractive index of 1.4. This value was based on both the manufacturer’s recommendation as well as prior literature on E. coli.13–15

Analysis

The typical loading and imaging process generated over 2500 images showing the evolution of a biofilm cross section throughout the experiment (some snapshots, Figure 3.2, numbered images). The features that could be captured were only limited by i) perfusion, which in most cases was negligible, ii) maximum fluid held by the reservoir, which was 70 ml, and iii) the thickness of the samples tested. Thicker samples would experience less stress at a given hydrostatic pressure, which meant the failure of a particular sample could not always be captured. A typical analysis protocol resulted in the stress-strain curve shown in OPTIONAL FIGURE BELOW.

OPTIONAL FIGURE. Cross sectional OCT images (top, numbered) acquired during the bulge test of a CE6 biofilm with subsequent 2D processing and stress-strain curve displayed below. Numbers with asterisks indicate data points corresponding to the images. Scale bar 500 µm. Biofilm did not fail during experiment.

The analysis protocol followed the assumptions of membrane theory, that: i) the thickness of the membrane is small in comparison to the in-plane dimensions ii) the bending stiffness is negligible and iii) due to (i) and (ii), in-plane stresses are assumed to be constant throughout the thickness of a membrane. Additionally, we assumed i) that there is an idealized equibiaxial deformation throughout the entire inflated membrane. In practice this is only true at the pole, and the deformation transitions to a constant width elongation at the clamped edge; also ii) that stress could be estimated using a spherical cap assumption, and strain could be estimated by changes in arc length of the deformed biofilm (Figure 3.3A).

Following the assumptions outlined above, the stress and strain states are as follows:

σij=σ1000σ20000

ϵij=ϵ1000ϵ2000ϵ3

We assume here that σ1=σ2=σ and ϵ1=ϵ2=ϵ across the entire membrane.

Applying Hooke’s law, the relationship between σ and ϵ is:

σ=E1-νϵ

Where E1-ν is the biaxial modulus of the material, referred to as the “modulus” in further text, and evaluated from the linear regions of each plot. We do not report the two quantities E and ν independently since they cannot be deconvolved using only one measurement method.

2D datasets were exported as tiff files and cropped using ImageJ. The resulting images were processed with in-house MATLAB scripts: the images were binarized; the top and bottom surfaces of the film were detected based on changes in pixel intensity and fit to fourth degree polynomials which were then used to estimate arc lengths of the top and bottom surfaces of the film (OPTIONAL FIGURE BELOW). Engineering strain ϵ was calculated as change in arc length divided by the original arc lengths of the top and bottom surfaces respectively.

Nominal stress was calculated with the equation for stress in a thin-walled spherical pressure vessel:

σ=PR2t

Where σ is film stress, P is the applied pressure, R is bulge radius of curvature, and t is film thickness. Further, R=a2+h2 where a is radius of aperture of the support disk and h is deflected height of the center of the film, directly measurable by OCT. For 2D datasets, the plotted data were averaged (top and bottom surfaces of the biofilm). True stress and strain were used for some films where indicated.

OPTIONAL FIGURE A) Spherical cap approximation for stress assumes that deformed biofilm is part of a larger spherical pressure vessel and the stress in the walls of the film balances the applied pressure. Strain is estimated as a difference in the arc length (red arrow) compared to the original (flat) length of the biofilm. B) Image processing scripts binarize and clean up OCT images and detect the top and bottom surfaces of the film over thousands of images (bottom right).

The ratio of aperture diameter 2a to film thickness t ranged from 16.6 – 25. Although the ideal ratio for the membrane assumption would be > 20, variations of stresses and strains through the thickness were neglected in our calculations based on the fact that the top and bottom surfaces of the biofilm did not show a significant difference in their stress-strain data (OPTIONAL FIGURE BELOW).

OPTIONAL FIGURE. We observed minimal discrepancy between stress-strain curves from the top vs. bottom surfaces of the biofilm.

CE6 vs E6 data

2D analysis revealed that the addition of the cysteine residue resulted in biofilms that were nearly three times stiffer (Figure 4.3) and rarely failed during our test. E6 was also ductile; showing yielding prior to failure.

Figure 4.3. CE6 (red) is three times stiffer than E6 (blue). Three biological replicates tested; select data shown. CE6 data point size two points smaller than E6 for clarity. Modulus based on fit to the initial linear region of each curve. Green asterisks indicate failure. CE6 did not fail during the tests above. N = 18; six biological replicates tested for E6. N = 6; four biological replicates tested for E6.

CE6 initial response is elastic and transitions to viscoelastic behavior over multiple loading cycles

Next, we investigated elasticity and changes in properties over multiple loading cycles. Since loading profiles were only limited by the syringe pump, we were able to apply a “sawtooth” loading pattern (Refer to figure above with the pressure profiles) and observe biofilm response simultaneously using OCT, processing the data in a similar manner.

Elasticity is the ability of a material to return to its original size and shape upon removal of applied loads.

A perfectly elastic material would return to its original size and shape, with either a linear or nonlinear stress-strain response that is independent of loading rate with superimposable stress-strain curves during loading and unloading. In a cyclic loading experiment, there is no phase lag between the applied stress and measured response.

A viscoelastic material, on the other hand, has both the characteristics of a solid and a fluid: properties are dependent on loading rate, with steeper response curves at faster loading rates. For a particular loading rate, the stress-strain response does not follow the same path for loading and unloading, and the area between the curves signifies energy dissipated during the process. Within a certain region, viscoelastic materials can completely regain their properties, while in (visco)plastic deformation, there is permanent deformation indicated by an x-intercept. In a cyclic loading experiment, a perfectly viscous material would have a phase lag of 90°, while a viscoelastic material would have a phase lag between 0 and 90° (Figure 3.11).

The bulge test was performed with a protocol similar to that of the “ramp” method, instead using a sawtooth profile (Figure 3.1, 3.12). The loading rate (slope) was 70 ml/min (maximum allowable by the syringe pump), to a maximum volume of 2 ml. The acquisition time was consistent at 19 ms. OCT images were processed with MATLAB with the same protocol described earlier. Inflection points were automatically detected in pressure vs. time (piecewise) plots and strain vs. time plots by identifying points with maximum change in mean and slope relative to surrounding data points, and then used to separate data points and plot information for each loading cycle. Representative results are shown below (N = 8, three biological replicates tested).

Figure 5.2. CE6 initially shows minimal lag between the applied load and response of the biofilm, increasing with number of cycles (A, pressure plotted as piecewise function in blue solid line scaled by 10-4 for ease of comparison; engineering strain as computed from OCT images as red data points; subset of cycles shown; additional inset in (B) magnified to highlight the increasing delay). Lag is consistent between biological replicates and increases almost linearly (B, 35 cycles shown; standard error).

CE6 showed minimal to no lag in the first loading cycle (Figure 5.2), with a correspondingly (nonlinear) elastic stress-strain trace (Figure 5.2A top left). The lag continued to increase over time with additional loading cycles, displaying viscoelastic behavior (Figures 5.2, 5.3).

Figure 5.3. First eight loading cycles for CE6, ordered top left to bottom right. CE6 displays an initially elastic response (top left), transitioning to a more viscoelastic response indicated by the increasing area between the loading and unloading curves.  Engineering strain axes are scaled by 103 for clarity.

E6 biofilms in contrast were challenging to test; however representative results are included below. Phase lag was significant from the first cycle (Figure 5.4) with identical loading conditions (resulting in roughly half the stress used for CE6, and a strain response that was ten times that of CE6). The film exhibited viscoplastic deformation from the first cycle (Figure 5.5). Ideally, loading conditions could have been adjusted to find an elastic regime; however repeated experiments were challenging due to the film properties.

Figure 5.4. E6 at similar conditions was not within the elastic regime and underwent permanent deformation.

Figure 5.5. E6 displays a viscoplastic response at pressures comparable to those used for CE6. The permanent deformation is mostly in the first two cycles; after which the response is viscoelastic.

TCEP: Data in folder; you can write the description and results since it’s not really related to mechanics.

Description of healing experiments

CE6 biofilms additionally expressing mWasabi were used for healing experiments (referred to in further text as “CE6mW”).

For self-healing experiments, multiple 3 mm punches from biofilms were separated from the polycarbonate as described above and then placed onto a fresh 2YT plate for healing.

A 25 µm thick TEM grid was used to make a cut through biofilms. Control films were placed onto fresh plates as well, without injury, for comparison.

Films were allowed to grow on the new plate at 37°C and monitored to assess the healing process.

Figure 7.1. A) Schematic of biofilm healing experiment. B) 3D printed petri dish holder fixed to OCT base plate.

OCT imaging of the cut and control biofilms was performed at 3, 6, and 16 hours. In order to image comparable regions each time, we 3D printed a petri dish holder (Figure 7.1B) and secured it to the OCT base plate, with a cavity that provided a snug fit for a single 60 mm diameter petri dish. The holder was at roughly a 10° slant to avoid noise from reflections in the OCT image at the air-biofilm interface. Only one biofilm punch was placed at the center of each petri dish, and permanent marks on both the petri dish and holder enabled the alignment of the petri dishes each time they were removed from the incubator for imaging. Although the vertical focus of the OCT could necessarily not be kept constant, we reproduced the scan settings in XY using the ThorImage OCT software to ensure that the scanned region was consistent for a given punch. Alignment was further confirmed using the camera images from the ThorImage OCT software, which provided an overall view of the sample itself as well as the specific scan box (Figures 7.2, 7.3, camera image insets with red rectangular boundaries).

3.4.2. Results: Quantifying the recovery of mechanical properties in healed biofilms

The OCT images of healing (Figures 7.2-7.3) provided a baseline for our bulge test time points. From the OCT images, the injured region showed clear healing in as soon as 3 hours; however the timescale of recovery of mechanical properties could only be assessed by the bulge test. Images were exported either as 3D renders at the same viewing angle and contrast (Figures 7.2-7.3, top), or as 2D cross sections at the same XZ plane in each image (Figures 7.2-7.3, bottom).

Figure 7.2. OCT scans of CE6mW healing on 2YT plate show rapid healing of injured biofilm. Top row exported as 3D renders. Scale: top, 4 x 4 mm scan box; bottom, 1 x 1 x 1 mm scan box. Insets are OCT camera images, manually cropped to region being scanned. Red rectangular outline represents scan box (automatic, from ThorImage OCT software).

Figure 7.3. OCT scans of CE6mW control on 2YT plate. Top row – patterned artifact due to software, not biofilm. Scale: top, 4 x 4 mm scan box; bottom, 1 x 1 x 1 mm scan box. Insets are OCT camera images, manually cropped to region being scanned. Red rectangular outline represents scan box (automatic, from ThorImage OCT software).

We attempted to perform bulge tests at 6, 12, 16, and 24 hour time points. Statistics are reported in Table 1, describing number of samples tested at each time point along with major points of failure. We discovered that healed biofilms could easily be separated from the nutrient plate by gently flooding the plate with PBS. After approximately 1-2 minutes, the punches freely separated from the agar underneath. Despite the gentle nature of this step, a few samples failed when peeled from the agar, with the previous injury reopening (second column, Table 1). The next major point of failure was during the initial pressure equilibration step (third column, Table 1). The final column in Table 1 detailed whether films failed during the bulge test itself.

| Table 1. Sample statistics for healed biofilms |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Total number of samples | Successfully peeled from agar and loaded (% of total) | Survived initial filling/pressure equilibration (% of total) | Failed within imposed pressures during bulge test (% of tested) |
| Original (Day 0) | 4 | 4 (100) | 4 (100) | 0 (0) |
| Controls (16 hours) | 4 | 4 (100) | 4 (100) | 2 (50) |
| Healed (6 hours) | 4 | 3 (75) | 0 (0) | - |
| Healed (12 hours) | 6 | 5 (83) | 4 (66) | 4 (100) |
| Healed (16 hours) | 7 | 6 (86) | 3 (43) | 3 (100) |
| Healed (24 hours) | 10 | 8 (80) | 0 (0) | - |

Properties of the immediately cut biofilm are not reported, since the cut film did not have the ability to maintain a pressure difference (OPTIONAL FIGURE). Fluid simply flowed through to the other side. This feature was also directly used in our experiments to determine the moment where the film failed, also visible under the OCT.

In comparing the various time points, we discovered that films healed for 6 hours healed to the extent that we could peel them from the agar and place them into our device; however the small differences in pressure resulting from the drop by drop filling of the device were sufficient to break the film once again.

Films healed for 12 and 16 hours could be readily tested. Surprisingly, 100% of films tested after 24 hours of healing could not withstand loading, suggesting that there is an optimal time frame for recovery of mechanical properties.

Next, we assessed the extent of recovery of mechanical properties using the 16 hour mark as our time point, comparing the original uncut film on day 0, the control (uninjured film grown on the healing plate), and healed films (Figure 7.4). It is worth noting that due to the increased thickness of the control and healed films relative to the original film (123 ± 2.78 and 120 ± 15.0 vs 103 ± 11.1 µm respectively, only considering films that were successfully tested), the range of stresses that could be imposed by our device (inversely proportional to the thickness of the film) was not directly comparable. However, we did observe a few clear differences among the films tested: first, the original (uninjured, day 7) films did not fail at the maximum pressure limits imposed by our device. The actual stress this corresponded to depended on the sample thickness and properties, but the highest observed stress tested on a CE6mW sample of 89 µm thickness was 7.21 kPa. While there was some variation in the range of applied stresses and observed strains during our test, we observed that CE6mW films on day 7 consistently tolerated stresses greater than 6.5 kPa (OPTIONAL FIGURE A).

In contrast, two out of four of the control films (16 hours) failed during our experiment. One failed at a stress of 4.68 kPa, while the other film failed at a stress of 3.65 kPa. (OPTIONAL FIGURE C)

100% of the healed films that could be loaded and tested failed within the range of stresses applied by our test (OPTIONAL FIGURE D). Given this result, we hypothesized that perhaps toughness (energy absorbed by the material prior to failure) could be used as an additional metric to assess the extent of healing (elaborated on in later text).

OPTIONAL FIGURE. Original, day 7 CE6mW films (A) do not fail within the limits of our test. Pictured maximum stress tested: 6.68 kPa. Cut films (B) immediately after injury cannot be tested, as the fluid freely flows through the tear (right). *50% of control films failed; pictured film (C) showed failure at stress: 4.68 kPa. All healed films failed (D); pictured maximum stress tolerated 3.32 kPa. Scale 200 µm.

From the stress-strain plots (Figure 7.4B), we observed that the modulus was similar across all three sets of films. This was surprising to note, but in order to assess whether it was a reasonable metric to compare, we analyzed two time points of healing (12 vs 16 hours), to assess which aspects of the film changed as the film healed.

Figure 7.4. A) Living materials can recover their mechanical properties after healing (stress-strain plots). B) Moduli for original, control (16 hours) and healed (16 hour) films are comparable: Original: 45.6 ± 6.8 kPa, Control: 36.5 ± 8.24 kPa, Healed: 41.4 ± 2.31 kPa. N = 3 for original and healed; N = 4 for control. Standard error reported. Individual biofilm failure not marked on stress-strain curves for clarity; see text for details.

In comparing biofilms healed for 12 vs. 16 hours, we found that the elastic moduli did not show a clear variation (12 hour: 52.7 ± 11.4; 16 hour: 41.4 ± 2.31 kPa). However, based on previous observations that every healed biofilm punch that could be tested also failed before the end of the test, we hypothesized that the toughness could be a quantity that varied with healing time. We estimated toughness by calculating the area under the stress-strain plots (Figure 7.5A, dotted lines connecting points with shaded area representing the toughness), and discovered that the healing time did indeed influence toughness – despite a large variability in the 16 hour samples, there was a clear increase in toughness of the healed films at the latter time point (Figure 7.5B).

Figure 7.4. A) Area under stress-strain plots of biofilms healed for 12 (orange) vs. 16 (green) hours represents toughness (B), which increases with time. N = 3. Standard error reported.

Cryomicrotome

A biopsy punch (Integra, 8 mm) was used to punch from the center of a biofilm (including the polycarbonate support underneath). The film was embedded in a glycol-based optimal cutting temperature compound (OCT, Sakura Finetek), flash frozen in liquid nitrogen, and left to equilibrate at -20°C overnight. Frozen serial sections (50 μm thick) were cut using a cryomicrotome (Tissue-Tek Cryo3, Sakura Finetek) at -20°C and mounted on coverslips for microscopy.

1.	Tsakalakos, T. & Hilliard, J. E. Elastic modulus in composition-modulated copper-nickel foils. J. Appl. Phys. 54, 734–737 (1983).

2.	Hsu, F. P. K., Liu, A. M. C., Downs, J., Rigamonti, D. & Humphrey, J. D. A Triplane Video-Based Experimental System for Studying Axisymmetrically Inflated Biomembranes. IEEE Trans. Biomed. Eng. 42, 442–450 (1995).

3.	Mohan, D. & Melvin, J. W. Failure properties of passive human aortic tissue. II-Biaxial tension tests. J. Biomech. 16, (1983).

4.	Cavinato, C. et al. Biaxial loading of arterial tissues with 3D in situ observations of adventitia fibrous microstructure: A method coupling multi-photon confocal microscopy and bulge inflation test. J. Mech. Behav. Biomed. Mater. 74, 488–498 (2017).

5.	Midgett, D. E., Quigley, H. A. & Nguyen, T. D. In vivo characterization of the deformation of the human optic nerve head using optical coherence tomography and digital volume correlation. Acta Biomater. 96, 385–399 (2019).

6.	Tonge, T. K., Atlan, L. S., Voo, L. M. & Nguyen, T. D. Full-field bulge test for planar anisotropic tissues: Part I-Experimental methods applied to human skin tissue. Acta Biomater. 9, 5913–5925 (2013).

7.	Wang, L., Tian, L., Huang, Y., Huang, Y. & Zheng, Y. Assessment of Corneal Biomechanical Properties with Inflation Test Using Optical Coherence Tomography. Ann. Biomed. Eng. 46, 247–256 (2018).

8.	Zamprogno, P. et al. Mechanical properties of soft biological membranes for organ-on-a-chip assessed by bulge test and AFM. ACS Biomater. Sci. Eng. 7, 2990–2997 (2021).

9.	O’Connell, P. A. & McKenna, G. B. Novel nanobubble inflation method for determining the viscoelastic properties of ultrathin polymer films. Rev. Sci. Instrum. 78, 14706 (2007).

10.	Turner, J. A., Menary, G. H. & Martin, P. J. Biaxial characterization of poly(ether-ether-ketone) for thermoforming: A comparison between bulge and in-plane biaxial testing. Polym. Eng. Sci. 59, 1853–1865 (2019).

11.	Marra, S. P., Kennedy, F. E., Kinkaid, J. N. & Fillinger, M. F. Elastic and rupture properties of porcine aortic tissue measured using inflation testing. Cardiovasc. Eng. 6, 123–131 (2006).

12.	Tabata, O., Kawahata, K., Sugiyama, S. & Igarashi, I. Mechanical property measurements of thin films using load-deflection of composite rectangular membranes. Sensors and Actuators 20, 135–141 (1989).

13.	Hart, S. J., Terray, A. V, Kuhn, K. L., Arnold, J. & Leski, T. A. Optical chromatography for biological separations. in Optical Trapping and Optical Micromanipulation vol. 5514 35 (SPIE, 2004).

14.	Liu, P. Y. et al. Real-time Measurement of Single Bacterium’s Refractive Index Using Optofluidic Immersion Refractometry. Procedia Eng. 87, 356–359 (2014).

15.	Balaev, A. E., Dvoretski, K. N. & Doubrovski, V. A. Refractive index of Escherichia coli cells. in Saratov Fall Meeting 2001: Optical Technologies in Biophysics and Medicine III vol. 4707 253–260 (SPIE, 2002).